## Slide 1
MEDICAL IMAGING · DEEP LEARNING
CT Brain Study Classifier
Training Flow & Model Architecture
MaxViT backbone + Gated-Attention MIL pooling + Grad-CAM++ explainability
Study-level, multi-slice, multi-window DICOM classifier — normal / near_normal / abnormal
Recipe adapted from Qari & Thafar (arXiv:2507.09630), “Brain Stroke Detection and Classification
Using CT Imaging with Transformer Models and Explainable AI” — extended to study-level MIL on DICOM.

## Slide 2
PIPELINE OVERVIEW
End-to-End Architecture
DICOM Series
studyID / seriesID bag of *.dcm slices
HU Windowing
3 clinical windows → 3-channel image
MaxViT Encoder
per-slice features (timm, 384px)
Gated-Attention MIL Pool
slices → 1 study embedding
Classifier Head
3-class logits + attention weights
Tensor flow: B bags × K slices × 3 × 384 × 384 → B×K×D features → B×D study vector → B×3 logits
1 · Multi-Window Input
• Raw pixels → Hounsfield Units (slope/intercept)
• Brain W80/L40 → channel 0
• Subdural W200/L80 → channel 1
• Bone W2800/L600 → channel 2
• Replaces grayscale-replicate with clinical windows
2 · Multiple-Instance Learning
• One label per STUDY, not per slice
• MaxViT encodes every slice independently
• Gated attention (Ilse 2018) pools K slices
• Attention weights → which slices mattered
• slice_chunk + grad-checkpoint bound memory
3 · Explainability (XAI)
• Grad-CAM++ on deep MaxViT conv layer
• stages.3.blocks.1.conv.conv2_kxk
• Heatmap = WHERE in the top-attended slice
• Attention bar charts = WHICH slices
• Logged to TensorBoard every N epochs
CT Brain Study Classifier · MaxViT + MIL + Grad-CAM++ · 3-class (normal / near_normal / abnormal)
2

## Slide 3
MODEL · models/maxvit_mil.py
MaxViT-MIL Network Architecture
Input bag
B × K × 3 × 384 × 384 (+ mask B×K)
Flatten slices
(B·K) × 3 × 384 × 384
MaxViT-tiny backbone (timm)
MBConv + block-attn + grid-attn → D-dim
Reshape + mask padding
B × K × D (padded slices zeroed)
Gated-Attention Pool
a = tanh(V·H) ⊙ σ(U·H); Z = Σ softmax(a)·H
Classifier head
LayerNorm → Dropout → Linear(D, 3)
Outputs
logits B×3 + attention B×K
Gated-Attention MIL Pooling (Ilse et al., 2018)
Per-slice score: aₖ = wᵀ ( tanh(V hₖ) ⊙ sigmoid(U hₖ) )
Masked softmax: αₖ = softmax(aₖ) over REAL slices only (padding → −∞)
Study embedding: Z = Σₖ αₖ · hₖ (attention-weighted sum)
attn_dim = 256. αₖ doubles as the interpretable per-slice importance.
Memory-bounded slice encoding
slice_chunk=96 → encode slices in chunks so activation memory is independent of bag size K.
grad_checkpoint=True (REQUIRED) → recompute backbone in backward; full fine-tune B16×K96×384
fits in ≈47.7 GB on an 80 GB A100 (OOMs without it). freeze_backbone option = head-only training.
Backbone is swappable (paper comparison)
maxvit_tiny_tf_384.in1k (default) · vit · tnt · convnext — via --backbone.
All loaded from timm as feature extractors (num_classes=0 → pooled embedding per slice).
CT Brain Study Classifier · MaxViT + MIL + Grad-CAM++ · 3-class (normal / near_normal / abnormal)
3

## Slide 4
TRAINING · train_main.py → engine/trainer.py
Training Flow
Build CSV
index every series folder path, slice_size, label
Study-grouped split
StratifiedGroupKFold 70/15/15 no patient leakage (assert disjoint)
Index studies
(sid, [slice paths], label) + inverse-freq class weights
StudyMILDataset
sort by InstanceNumber sample ≤96 slices → windows → aug
Length-bucketed batches
group similar bag sizes mil_collate → pad + attention mask
Forward pass
MaxViT-MIL → logits (AMP autocast, fp16)
Loss
weighted_ce / focal / cost_sensitive
Backward + step
GradScaler · clip 1.0 AdamW · grad-checkpoint
Cosine LR
warmup 2 ep → cosine decay
Validate / epoch
compute_metrics TensorBoard scalars
Checkpoint & early-stop
monitor = balanced_acc ↑ save best.pt / last.pt · patience 10
XAI logging (every N ep)
Grad-CAM++ overlays + MIL attention charts
Held-out TEST (best ckpt)
metrics + clinical operating point @ 0.95 not-normal sensitivity
↻ repeat up to 50 epochs (early stop on no improvement)
CT Brain Study Classifier · MaxViT + MIL + Grad-CAM++ · 3-class (normal / near_normal / abnormal)
4

## Slide 5
CONFIGURATION · config.py
Training Recipe & Clinical Decision Logic
Hyperparameters (defaults)
Backbone
maxvit_tiny_tf_384.in1k · pretrained
Input
384×384 · 3 HU-window channels
Bag size K
≤96 slices/study (random train / even eval)
Optimizer / LR
AdamW · 3e-4 · wd 1e-4 · cosine+warmup
Batch / epochs
16 studies · 50 (early stop 10)
Regularization
dropout 0.04 · label smoothing 0.05
Augmentation
hflip · rot ±15° · crop 0.85 · HU-jitter 0.05
Memory
grad-checkpoint · slice_chunk 96 · AMP fp16
Class weights
inverse-frequency (≈1 : 1.7 : 2.1)
Loss — --loss
weighted_ce (default) — CE × inverse-freq + smoothing
focal — (1−p)^γ CE for hard examples
cost_sensitive — expected-cost; under-calling abnormal→normal = 5×, near_normal→normal = 3×
Checkpoint selector — --monitor
balanced_acc (default) = mean per-class recall
Rewards recall on near_normal/abnormal; can't be gamed by over-calling.
Alternatives: f1 · not_normal_sensitivity · accuracy · auc
Clinical operating point (test only)
Score pathology as s = 1 − P(normal).
Pick threshold on VAL for ≥95% not-normal sensitivity,
apply to TEST → report op_test_sensitivity / specificity (no leakage).
“At 95% pathology sensitivity, here is the false-alarm rate.”
CT Brain Study Classifier · MaxViT + MIL + Grad-CAM++ · 3-class (normal / near_normal / abnormal)
5

## Slide 6
EVALUATION · engine/metrics.py
Metrics — What Each One Means
Every metric is logged per epoch to TensorBoard. Class index 0 = normal; “pathology” = near_normal or abnormal.
Metric
Plain-English meaning
How it is computed
accuracy
Fraction of studies whose top prediction is correct.
correct / total (argmax)
precision (macro)
When a class is predicted, how often it's right — averaged over classes.
mean per-class TP / (TP+FP)
recall (macro)
Of each class's true cases, how many were caught — averaged equally.
mean per-class TP / (TP+FN)
f1 (macro)
Balance of precision & recall, averaged over classes.
mean 2PR / (P+R)
balanced_acc ★
Mean recall across all 3 classes — the checkpoint selector.
mean( recall per class )
recall_<class>
Per-class sensitivity: recall_normal / _near_normal / _abnormal.
TP_class / (TP+FN)_class
not_normal_sensitivity
Of all true pathology, fraction flagged as NOT normal. (don't-miss)
pred≠normal & true≠normal / all path
normal_specificity
Of all true normals, fraction correctly called normal. (false-alarm)
pred=normal & true=normal / all normal
auc
Ranking quality, threshold-free. Macro one-vs-rest for 3 classes.
roc_auc_score(ovr, macro)
★ balanced_acc is the default --monitor: it rewards recall on every class and CANNOT be gamed by over-calling pathology (that would crush normal recall).
CT Brain Study Classifier · MaxViT + MIL + Grad-CAM++ · 3-class (normal / near_normal / abnormal)
6

## Slide 7
OPTIMIZATION · engine/losses.py
Loss Functions — Three Choices via --loss
weighted_ce · DEFAULT
Cross-entropy weighted by inverse class frequency, plus label smoothing 0.05.
L = Σ wᵧ · CE(logits, y)
wᵧ = N / (C · countᵧ)
Use when: imbalance is the main issue and error costs are symmetric.
focal · HARD EXAMPLES
Down-weights easy, confident cases so training focuses on hard ones.
L = (1 − pₜ)^γ · CE (γ = 2)
Risky here: amplifies the noisy normal ↔ near_normal boundary. Off by default.
cost_sensitive · CLINICAL COST
Minimizes EXPECTED cost under a cost matrix — missing pathology is expensive.
L = E_j[ pⱼ · C[y, j] ] + λ·CE
λ = 0.3 (stability / calibration)
Use when: under-calling pathology to normal must be punished hardest.
cost_sensitive — the cost matrix C[true, predicted]
predicted →
normal
near_norm
abnormal
normal
0
1
1
near_norm
3
0
1
abnormal
5
1
0
rows = true class, columns = predicted class
Reading the matrix
• Diagonal = 0 → correct predictions are free.
• Generic mistake = 1 → ordinary errors cost the same baseline.
• abnormal → normal = 5 → the worst possible miss (missed a real pathology).
• near_normal → normal = 3 → also heavily penalized.
Result: the model is pushed to keep probability mass OFF ‘normal’ for any case that might be pathological — exactly the clinical priority.
CT Brain Study Classifier · MaxViT + MIL + Grad-CAM++ · 3-class (normal / near_normal / abnormal)
7

## Slide 8
CLINICAL DESIGN
Decision Logic — Division of Labour
Three independent mechanisms each solve one part of “build a model that does not miss pathology.”
1 · LOSS
trains the right error profile
cost_sensitive makes under-calling
pathology→normal expensive.
→ shapes WHAT mistakes the model
is willing to make.
+
2 · MONITOR
picks the best model
balanced_acc = mean per-class recall
selects the best.pt checkpoint &
drives early stopping.
→ picks the best-DISCRIMINATING model.
+
3 · OPERATING POINT
sets the decision threshold
Tune threshold on val for ≥95%
pathology sensitivity, report on test.
→ sets the decision threshold to the
sensitivity clinicians require.
Operating point — step by step (no test leakage)
Score
s = 1 − P(normal) per study
Sweep val
ROC over thresholds on held-out val
Choose thr
max specificity s.t. sensitivity ≥ 0.95
Freeze
lock that single threshold
Apply to test
report op_test_ sensitivity / specificity
Why on val, applied to test? Choosing the threshold on one split and reporting on another removes optimistic bias —
it yields the defensible statement: “At 95% pathology sensitivity, the model's false-alarm rate on unseen data is X.”
CT Brain Study Classifier · MaxViT + MIL + Grad-CAM++ · 3-class (normal / near_normal / abnormal)
8